Format specification - rev2

Technical specification for encrypted EMG and IMU recording. This rev2 document updates the original design note dated 2026-04-23 to reflect the current implementation decisions and the accepted format simplifications.

# Revision summary

Status: This rev2 document is intended to replace the original prose design note for current implementation work. The reasoning below is normative where it changes parser or writer behavior.

| Area | Rev2 decision | Reasoning |
| --- | --- | --- |
| File structure | Remove the optional final summary record and define the file as header plus repeated batches only. | Recording can stop at any time because of power loss or deliberate shutdown. Batch trailers are the only reliable completion boundary. |
| Payload dispatch | Use payload_type as the sole chunk-layout discriminator and leave header byte 39 reserved. | This removes duplicated dispatch logic. Readers inspect one field to decide how to parse the payload. |
| Lead-off handling | Do not store a per-sample status byte. Store only a per-batch ch1_leadoff_samples summary in the control CSV. | This preserves useful signal-quality observability without increasing the binary payload by 1000 bytes per second. |
| Nonce scheme | Fix the nonce scheme to an entropy-seeded 8-byte session salt plus the 4-byte batch index. | The earlier RTC and session-id approach risks nonce reuse under a fixed key. Rev2 makes uniqueness independent of RTC state. |
| Cipher definition | State AES-256-GCM as the normative cipher for the current format. | The current STM32 firmware uses the hardware AES path, so the specification should describe the implemented algorithm rather than alternatives. |
| Current filenames | Document the current SNNNNNNN.EMX and SNNNNNNN.CSV filenames while noting that longer names remain a future option. | The emitted files must match FatFS 8.3 constraints today, but naming does not change the on-disk byte format. |

# 1. Purpose and scope

The recording format is intended for efficient and secure storage of EMG and IMU data acquired by the AstroMoWe embedded device. The format is optimized for real-time logging on STM32L462-class hardware while remaining simple to parse on desktop tools.

Rev2 defines the current implemented format and removes design-note ambiguities that are no longer useful for interoperable readers and writers.

# 2. Output files

Each recording session produces two coordinated files that share the same session identifier.

| File | Current implementation | Purpose |
| --- | --- | --- |
| Primary binary file | SNNNNNNN.EMX | Encrypted measurement data |
| Secondary control file | SNNNNNNN.CSV | Plaintext operational metadata |

Naming note: The current firmware uses 8.3 filenames because FatFS long filename support is disabled. Longer SESSION_NNNNNNN names remain an optional future change and do not affect the byte format.

# 3. File layout

A recording file is laid out as a 64-byte plaintext file header followed by a repeated sequence of plaintext batch header, encrypted payload, and plaintext trailer.

Rev2 change: The optional final summary record from the original design note is removed. Completed batches are the only durable recovery boundary.

# 4. File header

| Offset | Size | Field | Rev2 value or meaning |
| --- | --- | --- | --- |
| 0 | 4 | magic | "EMGX" |
| 4 | 1 | version | 1 |
| 5 | 1 | header_size | 64 |
| 6 | 1 | payload_type | 2 = EMG ch1 + IMU chunk |
| 7 | 1 | batch_duration_s | 10 |
| 8 | 4 | session_id | u32, matches the filename |
| 12 | 12 | device_uid | STM32 96-bit unique device ID |
| 24 | 6 | start_time_bcd | RTC BCD: sec,min,hr,date,mon,yr |
| 30 | 2 | emg_rate_hz | 1000 |
| 32 | 2 | imu_rate_hz | 100 |
| 34 | 1 | cipher_id | 1 = AES-256-GCM |
| 35 | 1 | key_id | From recording.key |
| 36 | 1 | key_version | From recording.key |
| 37 | 1 | nonce_scheme | 1 = entropy-seeded random salt + batch counter |
| 38 | 1 | emg_pga_gain | ADS1292R CH1 PGA gain (4) |
| 39 | 1 | reserved | 0 |
| 40 | 8 | nonce_prefix | Per-session nonce salt |
| 48 | 10 | ads_regs | ADS1292R snapshot: ID, CONFIG1, CONFIG2, LOFF, CH1SET, CH2SET, RLDSENS, LOFFSENS, RESP1, RESP2 |
| 58 | 2 | emg_ref_mv | 2420 |
| 60 | 4 | reserved | 0 |

# 5. Batch header

Each batch begins with a 48-byte plaintext header. The header is written up front with nominal full-batch counts and may be rewritten with the actual counts if recording stops early.

| Offset | Size | Field | Meaning |
| --- | --- | --- | --- |
| 0 | 4 | marker | "BATC" |
| 4 | 4 | batch_index | u32, from 0 |
| 8 | 6 | start_time_bcd | RTC BCD at batch start |
| 14 | 1 | batch_duration_s | 10 |
| 15 | 1 | reserved | 0 |
| 16 | 4 | emg_sample_count | u32 |
| 20 | 4 | imu_sample_count | u32 |
| 24 | 4 | plaintext_payload_bytes | u32 |
| 28 | 4 | ciphertext_payload_bytes | u32, equal to plaintext size |
| 32 | 12 | nonce | nonce_prefix (8) plus batch_index as u32 little-endian (4) |
| 44 | 4 | reserved | 0 |

# 6. Payload definition

Payload dispatch: payload_type is the only chunk-layout discriminator in rev2. The current implemented payload_type is 2, meaning one-second chunks of EMG ch1 plus IMU.

Chunk size: One chunk equals 6400 bytes: int32 ch1[1000] plus ImuSample[100]. A full 10-second batch contains 10 chunks and therefore 64000 plaintext bytes.

| Section | Type | Count | Bytes | Meaning |
| --- | --- | --- | --- | --- |
| ch1 | int32 LE | 1000 | 4000 | ADS1292R channel 1, 24-bit signed value sign-extended to int32 |
| imu | ImuSample | 100 | 2400 | 100 Hz IMU slots derived from the EMG sample clock |

Lead-off policy: Per-sample lead-off status is not stored in the binary payload. Instead, the firmware counts CH1 lead-off samples and reports the per-batch summary in the control CSV as ch1_leadoff_samples.

## 6.1 ImuSample layout

| Offset | Type | Fields |
| --- | --- | --- |
| 0 | int16 x3 | ax, ay, az |
| 6 | int16 x3 | gx, gy, gz |
| 12 | uint32 x3 | mx, my, mz |

# 7. Encryption and nonce

Cipher: AES-256-GCM is the normative cipher for this format. Each batch is encrypted independently as one GCM stream over its plaintext payload.

Nonce: The 96-bit batch nonce is nonce_prefix (8 bytes) plus batch_index encoded as a 32-bit little-endian value.

Rev2 nonce rationale: nonce_prefix is a per-session entropy salt derived from ADS1292R analog noise, the free-running CPU cycle counter, and the device UID. This avoids dependence on RTC state or monotonic session numbers and therefore avoids nonce reuse risk under a fixed key.

# 8. Batch trailer

| Offset | Size | Field | Meaning |
| --- | --- | --- | --- |
| 0 | 4 | crc32_plaintext | CRC32 of the plaintext payload |
| 4 | 16 | tag | AES-GCM authentication tag |
| 20 | 4 | end_marker | "ENDB" |

Recovery rule: A batch missing the trailer or ENDB marker is treated as incomplete and discarded by readers. Scanning resumes at the next BATC marker.

# 9. Secondary control CSV

The control CSV contains one row per completed batch and no raw EMG or IMU measurement values.

| Column | Meaning |
| --- | --- |
| session_id | Session identifier, matches the filename |
| batch_index | Zero-based batch counter |
| start_timestamp | Batch start time in BCD YYMMDDhhmmss |
| end_timestamp | Batch end time in BCD YYMMDDhhmmss |
| emg_sample_count | Nominally 10000 for a full 10-second batch |
| imu_sample_count | Nominally 1000 for a full 10-second batch |
| unencrypted_payload_bytes | Nominally 64000 for a full batch |
| encrypted_payload_bytes | Same size as plaintext payload |
| write_time_ms | Time spent in f_write and f_sync for the batch |
| crc32_plaintext | CRC32 of the plaintext payload |
| dropped_samples | Samples lost because the acquisition ring overflowed |
| temperature_c | Optional IMU temperature field |
| error_flags | Hex bit field: WRITE, SYNC, RTC, DROPPED, AES, PARTIAL |
| storage_status | OK or ERR |
| ch1_leadoff_samples | Count of CH1 lead-off samples in the batch |

# 10. Physical scaling

EMG scaling: emg_uV = ch1 * emg_ref_mv * 1000.0 / ((1 << 23) * emg_pga_gain). With emg_ref_mv = 2420 and emg_pga_gain = 4, one count is approximately 0.0721 microvolts.

Accelerometer: g = raw / 16384 for the +/-2 g range.

Gyroscope: dps = raw / 131 for the +/-250 dps range.

Magnetometer: mx, my, and mz are stored as 18-bit unsigned values inside uint32 fields. A null field is 131072.

# 11. Compatibility and versioning

Current reader rule: Readers must require version == 1 and payload_type == 2 for this rev2 document. Anything else is outside the current implemented format.

Legacy note: Earlier development recordings are obsolete and incompatible. This includes pre-BATC files and early raw-status recordings that reused payload_type 2 with a different chunk size and a different nonce derivation.

Future changes: If a future payload layout changes, payload_type should be bumped. If the header semantics change, version should be bumped.

# 12. References

This rev2 document aligns with the current implementation sources: the firmware writer in recording.c, acquisition support in acquisition.c, the desktop decoder decode_emgx.py, and the aligned markdown specification FORMAT.md.